## Slide 1
AT-TPC
Gas Handling System
Zach Meisel

## Slide 2: Science to Address
Rare processes that require high detection efficiency and large acceptance
Low energy processes that are traditionally difficult to measure due to the short range of the reaction products in matter
Several target-compositions desirable, many gaseous

## Slide 3: …what’s the optimal detector system?
Given:
  Rare processes that require high detection efficiency and large acceptance
  Low energy processes that are traditionally difficult to measure due to the short range of the reaction products in matter
  Several target-compositions desirable, many gaseous
Active-Target
Time-Projection
Chamber
Van Wizard

## Slide 4: AT-TPC: Basic Design
Beam
Active Target = High(ish) Pressure Gas-Cell
Time-Projection Chamber = Gas-Cell w/ voltage gradient with e- detectors at one end
Surround with Magnetic field to effectively extend chamber size & aid with PID

## Slide 5: AT-TPC: The Gas Matters!
Initial e- created proportional to energy loss (i.e. Z2 and E-1)
Secondary e- created due to acceleration voltage
  Multiplication factor, g,
  where the Townsend coefficient, α,
  and ε is the electric field, p is the pressure, and A,B are gas-specific constants
Gain is very sensitive to gas properties!

## Slide 6: Gas Handling Requirements
Must maintain gas composition
  To within 0.03% for gain variation of 1%! (D.Suzuki)
Must maintain gas pressure
  To within 0.1% for gain variation of 1%! (D.Suzuki)
Must have low-rate of gas loss
  Only have 100std.L of 3He and more is difficult to get
  i.e. require high leak-tightness
Flow
the gas
Circulate the gas

## Slide 7: Gas Handling in a Nutshell
Target Volume
Compression
Purification

## Slide 8
SGC
ATGC
SP1
BV1
SP2
BV2
General Gas-handling Design
M
M
PT
PT
DPT
PT1
PT2
DPT1
TGC2
MV2
MV1
Reg
RG1
RG2
M
M
MFC2
MFC1
GA1
Exhaust
Exhaust
MV6
C1
MDV1
CV1
P1
P2
F2
F1
MFC4
PT3
ST1
CT1
HE1
M
TGC1
MFC3
MV7
MV3
MV4
N2
MDV2
MV5
MS1
MB1
ATGC: Active Target Gas Cylinder
BT: Ballast/Buffer Tank
BV: Butterfly Valve
C: Compressor
CT: CO2 Cold-Trap
CV: Check Valve
DPT: Differential Pressure Transducer
F: Filter
GA: Gas Analyzer
HE: Heat Exchanger
MB: Mixing Baffle
MDV: Motor-Driven Valve
MFC: Mass Flow Controller
MS: Molecular Sieve
MV: Manual Valve
P: Purifier
PT: Pressure Transducer
RG: Regulator
SGC: Shield Gas Cylinder
SP: Scroll Pump
ST: Storage Tank
TGC: Target Gas Cylinder
Used in CO2 case
Used in non-circulating case
Used to check mixing directly
Break points for non-circulating
MV8
DPT
DPT2
MDV5
MV9
MV10
CT2
MDV3
MDV4
BT1
Reg
TGC3
RG3
MV11
MV12
Reg
RG4
TGC4
Reg
Target Volume
Compression
Purification

## Slide 9: Pressure Control

## Slide 10: Pressure Control
Pressure Gauge
“Down-stream” control: Fix in-flow, adjust out-flow
Gas Volume
Gas Source
Fixed
Flow
Restrictor
Adjustable
Flow
Restrictor
Pumping
Feedback

## Slide 12: Pressure Control Data
Monitor pressure stability around set-point
  For gases: 4He, Ar, N2
  At pressures (Torr): 25, 40, 80, 200, 500

## Slide 13: Pressure Control Results
From ATTPC gas-handling requirements document:
“The pressure stability of the system in the active gas target chamber is also extremely important. The expected range of operating pressure is from 5 to 760 torr. Numerical simulations show that the total pressure should be maintained within 0.1% to obtain an uncertainty in the signal gain of 1% [Suzuki]. such uncertainty corresponds to the expected resolution of the micromegas. Pressure and temperature control systems should be able to achieve this level of stability.”
Gas flow 5 liters per minute
(±0.5 for N2 and Ar, ±1 for He)
Adequate pressure control down to 80Torr

## Slide 14
Compression

## Slide 15
Compression
Primary Purpose: Overcome flow-impedance from purifier
….also remaining system conductance:
15psig = 776Torr above atmospheric pressure
40 kPa = 300Torr

## Slide 16
Compression Test
Flow Meter
Pressure Gauge
Gas Volume
Gas Source
Fixed
Flow
Restrictor
Adjustable
Flow
Restrictor
Pump
Feedback
Compressor
Tube
Obtain compression (i.e. post-compressor pressure) via Hagen-Pouiselle Equation:

## Slide 17
Buffer Tank

## Slide 18
Buffer Tank Volume
Say dP=1%*550Torr, Vsyst=500LPbuff=760torr  Vbuff=3.6L
(…more conservative (almost unreasonable) estimate gives 11L)

## Slide 19
Purification

## Slide 20
Remove impurities via purifier/filter
  Zeolite trap w/ a mesh
  Purifier= remove via molecular selection
  Filter = remove based on size via mesh
Remove purifier poisons via cold trap
  E.g. 4He+CO2 (target + quench gas)
  Must remove old CO2 and add new CO2 each cycle
Heat gas back up via heat-exchanger
Purification
Quench Gas: Absorbs ions & photons released in ionization process which would produce delayed e-

## Slide 21
“Layout Grid”
Purification + Mass Flow Control
Surface-Mount technology via Swagelok
  Minimize volume necessary
  High leak-tightness
  Drawn via Swagelok IGC-II Configurator
  Miscellaneous parts required & instructions provided by Swagelok

## Slide 22
Purity Analysis

## Slide 23
Purity Analysis
Monitor target-quench gas concentration
Monitor contaminants
  O2 and H2O attenuate e- drift
  Must be kept below 10-20ppm concentration for 1% σVdrift
Must be able to detect several gases from the 10’s of % to ppm level!
Always below required 0.03% fluctuation

## Slide 24
Quadrupole Mass Analyzer
Purity Analysis
Must be able to detect several gases from the 10’s of % to ppm level
Van Wizard

## Slide 25
System Control
Must integrate I/O between several kinds of devices:
National Instruments Compact RIO controlled via LabView

## Slide 26: ...still lots to be done
SGC
ATGC
SP1
BV1
SP2
BV2
General Gas-handling Design
M
M
PT
PT
DPT
PT1
PT2
DPT1
TGC2
MV2
MV1
Reg
RG1
RG2
M
M
MFC2
MFC1
GA1
Exhaust
Exhaust
MV6
C1
MDV1
CV1
P1
P2
F2
F1
MFC4
PT3
ST1
CT1
HE1
M
TGC1
MFC3
MV7
MV3
MV4
N2
MDV2
MV5
MS1
MB1
ATGC: Active Target Gas Cylinder
BT: Ballast/Buffer Tank
BV: Butterfly Valve
C: Compressor
CT: CO2 Cold-Trap
CV: Check Valve
DPT: Differential Pressure Transducer
F: Filter
GA: Gas Analyzer
HE: Heat Exchanger
MB: Mixing Baffle
MDV: Motor-Driven Valve
MFC: Mass Flow Controller
MS: Molecular Sieve
MV: Manual Valve
P: Purifier
PT: Pressure Transducer
RG: Regulator
SGC: Shield Gas Cylinder
SP: Scroll Pump
ST: Storage Tank
TGC: Target Gas Cylinder
Used in CO2 case
Used in non-circulating case
Used to check mixing directly
Break points for non-circulating
MV8
DPT
DPT2
MDV5
MV9
MV10
CT2
MDV3
MDV4
BT1
Reg
TGC3
RG3
MV11
MV12
Reg
RG4
TGC4
Reg

## Slide 27: Thanks to:
AT-TPC Group (…probably missing some people here)
  Tan Ahn, Faisal Abu-Nimeh, Daniel Bazin, Saul Beceiro-Novo, Abigail Bickley, Zbigniew Chajecki, Adam Fritsch, Ian La Valley, Bill Lynch, Wolfgang Mittig, Aimee Shore, Nathan Usher, John Yurkon
Others
  Steve Bricker, Fernando Montes, Pat Mussel, Dave Sanderson, Hendrik Schatz, Andy Thulin
For more, check-out the slightly out-of-date (but hopefully soon to be updated) wiki page:
https://groups.nscl.msu.edu/tpc/wiki/doku.php?id=gas_handling

## Slide 28: SUPPLEMENTAL SLIDES

## Slide 29: AT-TPC: General Design

## Slide 30: MICROMEGAS

## Slide 31: Bethe-Formula

## Slide 32
Basic Layout: 3-Level Rack
Top Level: Surface-Mount Components
Mass Flow Controllers
Purifier
Heat-exchanger & Cold-trap Ports
Assorted Valves
Middle Level: Buffer Tanks
Small buffer tank for recirculation
Large buffer tank for initial system pressurization
Assorted Valves
Bottom Level: Pumping & Compression
Port leading to storage tank
Compressor
Scroll-Pumps
Butterfly valves
Assorted Valves
On the side:
Storage tank
Gas Cylinders
Purity Analyzer